Human Alpha-1 antitrypsin    (A1AT) immunoassay kit

Catalogue Number: 31390

For the quantitative determination of human Alpha-1 antitrypsin

concentrations in serum and plasma samples

This package insert must be read in its entirety before using this product.

Website: www.immunodiagnostics.com.hk

E-mail: info@immunodiagnostics.com.hk

Tel: (+852) 3502 2780

Fax: (+852) 3502 2781

TABLE OF CONTENTS

| Contents | Page |
| --- | --- |
| INTRODUCTION | 1 |
| PRINCIPLE OF THE ASSAY | 1 |
| INTENDED USE | 1 |
| REAGENTS SUPPLIED | 2 |
| OTHER MATERIALS REQUIRED, BUT NOT PROVIDED | 2 |
| STORAGE | 2 |
| PREPARATION OF REAGENTS | 3 |
| PREPARATION OF STANDRADS, CONTROL AND SAMPLE | 4 |
| ASSAY PROCEDURE | 5 |
| CALCULATION | 5 |
| TYPICAL STANDARD CURVE | 6 |
| REFERENCES | 7 |
| SUMMARY OF ASSAY PROCEDURE | 8 |

INTRODUCTION

Alpha-1 antitrypsin (A1AT), also known as alpha-1 protease inhibitor or SERPINA1, is a 52 kDa serine protease inhibitor, which is mainly synthesized in the liver, but also produced in monocytes, macrophages, dendritic cells, pulmonary alveolar cells, intestinal and corneal epithelium [1]. The primary function of A1AT is to inhibit the actions of proteolytic enzymes, such as neutrophil elastase (NE), proteinase 3 (PR3) and cathepsin G (CG), and to provide essential protection to host tissues from non-specific injury during periods of inflammation [2]. A1AT has also been reported to regulate neutrophil chemotaxis, to enhance insulin secretion and protect β-cells against cytokine-induced apoptosis, and to possess the anti-inflammatory as well as immunomodulatory properties [3-5]. Systemic deficiency of A1AT due to genetic mutations can result in a number of diseases, such as chronic obstructive pulmonary disease (COPD), systemic sclerosis, liver injury, cirrhosis and hepatocellular carcinoma [6-8]. Furthermore, A1AT disorder has also been shown to be involved in the development of diabetes mellitus [9, 10].

PRINCIPLE OF THE ASSAY

This assay is a quantitative sandwich ELISA. The immunoplate is pre-coated with a polyclonal antibody specific for human A1AT. Standards and samples are pipetted into the wells and any human A1AT present is bound by the immobilized antibody. After washing away any unbound substances, a horseradish peroxidase (HRP)-linked polyclonal antibody specific for human A1AT is added to the wells. After a final wash step, an HRP substrate solution is added and colour develops in proportion to the amount of human A1AT bound initially.  The assay is stopped and the optical density of the wells determined using a microplate reader. Since the increases in absorbance are directly proportional to the amount of captured human A1AT, the unknown sample concentration can be interpolated from a reference curve included in each assay.

INTENDED USE

This human A1AT ELISA kit is designed for quantification of human A1AT in serum and plasma samples.

REAGENTS SUPPLIED

Each kit is sufficient for one 96-well plate and contains the following components:

1. Micro-titre Strips (96 wells)-Coated with a polyclonal antibody

against human A1AT, sealed.

2. 10×Wash buffer-40 ml.

3. 5×Assay buffer-60 ml.

4. 100×Detection antibody solution-A HRP labelled polyclonal antibody

against human A1AT, 0.12 ml.

5. Human A1AT standard-20 ng of native human A1AT in a buffered

protein base,  lyophilised.

6. Substrate solution- 12 ml, ready for use.

7. Stop solution- 12 ml, ready for use.

OTHER MATERIALS REQUIRED, BUT NOT PROVIDED

1. Pipettes and pipette tips.

2. 96-well plate or manual strip washer.

3. Buffer and reagent reservoirs.

4. Paper towels or absorbent paper.

5. Plate reader capable of reading absorbency at 450 nm.

6. Distilled water or deionized water.

STORAGE

The kit should be stored at 2-8°C upon receipt, and all reagents should be equilibrated to room temperature before use. Remove any unused antibody-coated strips from the human A1AT microplate, return them to the foil pouch and re-seal. Once opened, the strips may be stored at 2-8°C for up to one month.

PREPARATION OF REAGENTS

Bring all reagents and materials to room temperature before assay.

A. 1×Assay buffer.

Prepare 1×Assay buffer by mixing the 5×Assay buffer (60 ml) with 240

ml of distilled water or deionized water.  If precipitates are observed in

the 5× Assay  buffer  bottle, warm  the  bottle  in  a  37°C water  bath

until  the precipitates disappear.  The 1×Assay buffer may be stored at

2-8°C for up to one month.

B. 1×Wash buffer.

Prepare 1×Wash buffer by mixing the 10×Wash buffer (40 ml) with 360

ml of distilled water or deionized water.  If precipitates are observed in

the 10× Wash  buffer  bottle, warm  the  bottle  in  a  37°C  water

bath until  the  precipitates disappear.  The 1×Wash buffer may be

stored at 2-8°C for up to one month.

C. 1×Detection antibody solution.

Spin down the 100×Detection antibody solution briefly and dilute the

desired amount of the antibody 1:100 with 1×Assay buffer, 100 (l  of

the 1×Detection antibody solution is required per well.  Prepare only

as much 1×Detection antibody solution as needed.  Return the

100×Detection antibody solution to 2-8°C immediately after the

necessary volume is removed.

PREPARATION OF STANDRADS, CONTROL AND SAMPLE

Human A1AT standards: Reconstitute the lyophilised standard with 1 ml of

1×Assay buffer to generate a standard stock solution of  20 ng/ml.  Allow the

standard to sit for 10 minutes with gentle agitation prior to making dilutions.

Prepare serially diluted standards using 1×Assay buffer as follows:

| Standard volume | Volume of 1×Assay buffer | Concentration |
| --- | --- | --- |
| 20 ng/ml stock | - | 20 ng/ml |
| 250 (l of 20 ng/ml | 250 (l | 10 ng/ml |
| 250 (l of 10 ng/ml | 250 (l | 5 ng/ml |
| 250 (l of 5 ng/ml | 250 (l | 2.5 ng/ml |
| 250 (l of 2.5 ng/ml | 250 (l | 1.25 ng/ml |
| 250 (l of 1.25ng/ml | 250 (l | 0.625ng/ml |
| 250 (l of 0.625 ng/ml | 250 (l | 0.312 ng/ml |

1×Assay buffer serves as the zero standard (0 pg/ml).

Note: The reconstituted standard stock should be aliquoted and stored at -80ºC

for up to one month.   Avoid repeating freezing/thawing cycles.  Please do not

store the diluted standard solutions.

Sample preparation

Serum or plasma sample is generally required a 1000000-fold dilution in this

assay.  A three-step dilution is suggested. Step1- Add 10 µl of sample to 990

µl of 1× Assay buffer to give a 100-fold diluted sample solution.  Step 2- Add

5 µl of the 100-fold diluted sample solution to 495 µl of 1×Assay buffer to

give a 10000-fold diluted sample solution.  Step 3- Add 5 µl of the

10000-fold diluted sample solution to 495 µl of 1×Assay buffer to give a final

1000000-fold diluted sample solution.  If a sample has A1AT level greater than the highest standard, the sample should be diluted further and the assay should be repeated.

Assay procedure

It is recommended that all standards, control and samples should be assayed

in duplicate.

1. Add 100 µl of standard or sample per well,   incubate at room

temperature for 1 hour.

2. Discard the content and tap the plate on a clean paper towel to remove

residual solution in each well.    Add 300 µl of 1×Wash buffer to each

well and incubate for 1 minute.  Discard the 1×Wash buffer and tap the

plate on a clean paper towel to remove residual wash buffer.  Repeat the

wash step for a total 3 washes.

3. Add 100 µl of 1×Detection antibody solution to each well, incubate at

room temperature for 1 hour.

4. Wash each well 4 times as described in step 2.

5. Add 100 µl of Substrate solution to each well, incubate at room

temperature for 15 minutes.  Protect from light.

6. Add 100 µl of Stop solution to each well, gently tap the plate frame for a

few seconds to ensure thorough mixing.

7. Measure absorbance of each well at 450 nm immediately.

Calculation

1. Subtract absorbance of  blank from that of standards and samples.

2. Generate a standard curve by plotting the absorbance obtained (y-axis)

against human A1AT concentrations (x-axis).  The best fit line can be

generated with any curve-fitting software by regression analysis.  Any

curve of 4-parameter or log-log curve fitting can be used for

calculation.

3. Determine human A1AT concentration of samples from standard curve

and multiply the value by the dilution factor.

Typical standard curve

The following standard curve is provided for demonstration only. A standard curve should be generated for each set of sample assay.

| Human A1AT (ng/ml) | Absorbance (450 nm) | Blanked Absorbance |
| --- | --- | --- |
| 0 | 0.08 | 0 |
| 0.312 | 0.129 | 0.049 |
| 0.625 | 0.166 | 0.086 |
| 1.25 | 0.267 | 0.187 |
| 2.5 | 0.445 | 0.365 |
| 5 | 0.816 | 0.736 |
| 10 | 1.47 | 1.39 |
| 20 | 2.55 | 2.47 |

REFERENCE

1.	Janciauskiene SM, Bals R, Koczulla R, Vogelmeier C, Kohnlein T, Welte T: The discovery of alpha1-antitrypsin and its role in health and disease. Respir Med 2011, 105:1129-1139.

2.	Miravitlles M: Alpha-1-antitrypsin and other proteinase inhibitors. Curr Opin Pharmacol 2012, 12:309-314.

3.	Bergin DA, Reeves EP, Meleady P, Henry M, McElvaney OJ, Carroll TP, Condron C, Chotirmall SH, Clynes M, O'Neill SJ, McElvaney NG: alpha-1 Antitrypsin regulates human neutrophil chemotaxis induced by soluble immune complexes and IL-8. J Clin Invest 2010, 120:4236-4250.

4.	Kalis M, Kumar R, Janciauskiene S, Salehi A, Cilio CM: alpha 1-antitrypsin enhances insulin secretion and prevents cytokine-mediated apoptosis in pancreatic beta-cells. Islets 2010, 2:185-189.

5.	Jonigk D, Al-Omari M, Maegel L, Muller M, Izykowski N, Hong J, Hong K, Kim SH, Dorsch M, Mahadeva R, et al: Anti-inflammatory and immunomodulatory properties of alpha1-antitrypsin without inhibition of elastase. Proc Natl Acad Sci U S A 2013, 110:15007-15012.

6.	Brebner JA, Stockley RA: Recent advances in alpha-1-antitrypsin deficiency-related lung disease. Expert Rev Respir Med 2013, 7:213-229; quiz 230.

7.	Barnes TC, Cross A, Anderson ME, Edwards SW, Moots RJ: Relative alpha(1)-anti-trypsin deficiency in systemic sclerosis. Rheumatology (Oxford) 2011, 50:1373-1378.

8.	Teckman JH: Liver disease in alpha-1 antitrypsin deficiency: current understanding and future therapy. COPD 2013, 10 Suppl 1:35-43.

9.	Mansuy-Aubert V, Zhou QL, Xie X, Gong Z, Huang JY, Khan AR, Aubert G, Candelaria K, Thomas S, Shin DJ, et al: Imbalance between neutrophil elastase and its inhibitor alpha1-antitrypsin in obesity alters insulin sensitivity, inflammation, and energy expenditure. Cell Metab 2013, 17:534-548.

10.	Wang Y, Xiao Y, Zhong L, Ye D, Zhang J, Tu Y, Bornstein SR, Zhou Z, Lam KS, Xu A: Increased Neutrophil Elastase and Proteinase 3 and Augmented NETosis Are Closely Associated with beta-cell Autoimmunity in Patients with Type 1 Diabetes. Diabetes 2014.

SUMmary OF Assay procedure

Add 100 μl of Standard, control or sample to each well. ( Incubate at room temperature for 1 hour. ( Aspirate and wash each well three times. ( Add 100 μl of 1×Detection antibody solution to each well. ( Incubate at room temperature for 1 hour. ( Aspirate and wash each well four times. ( Add 100 μl of Substrate solution to each well. ( Incubate at room temperature for 15 minutes. ( Add 100 μl of Stop solution to each well. ( Measure absorbance of each well at 450 nm. ( Calculation

ImmunoDiagnostics Limited

5F, Biotech Centre 2 (11W), No. 11 Science Park

West Avenue, Hong Kong Science Park, Shatin. NT, Hong Kong

Website: www.immunodiagnostics.com.hk

E-mail: sales@immunodiagnostics.com.hk

Tel: (+852) 3502 2780

Fax: (+852) 3502 2781

# Text boxes

Human A1AT standard curve (4-parameter)